# KHAMSIN	User manual

V14-AUGUST 2026-PCB

# This manual must be read

before using the KHAMSIN Automated Cleaning System for flexible endoscopes.

This ensures that it is used under the safest and most efficient conditions.

Partners for Endoscopy Ltd (PFE Medical), Hartley House, Galveston Grove, Fenton, Stoke on Trent, UK

Advena Limited, Tower Business Centre, 2nd Floor, Tower Street, Swatar, BKR 4013, Malta

For any questions concerning the equipment, its maintenance, breakdowns, or end-of-life, please contact your after-sales representative.

The description of the automated cleaning system is accurate on the date this manual is printed.

Table of contents

- Introduction	4

  - Presentation of the KHAMSIN	4

    - Intended use	4

    - Certified endoscopes	4

    - Users and environment	4

    - Training	4

  - Recommended use	4

  - Accessories	4

  - Technology	4

  - Certification	5

  - Technical characteristics	5

- Regulatory information, safety rules, and operating precautions	6

  - Safety rules	6

    - Handling	6

    - Environmental conditions	6

    - Use	6

    - Electricity	6

    - Electromagnetism	7

    - Upkeep and maintenance	7

  - Marking	8

- Installation, adjustments, and commissioning	9

  - Packaging contents	9

  - Presentation of the interfaces	10

  - Installation	10

    - Verification	10

    - Placement	10

    - Connection to electrical power	10

    - Connection of accessories	10

- Use	11

  - Switching on	11

  - Switching off	11

  - Prerequisites to endoscope disinfection	11

  - Cleaning an endoscope	12

- Upkeep	12

  - External cleaning	12

  - Cleaning of the chamber	12

- Scheduled maintenance	13

- Troubleshooting	13

- Transport and storage	13

- Disposal	13

- Warranty	13

- Contact	14

## INTRODUCTION

# Presentation of the KHAMSIN

### Intended use

The KHAMSIN is an Automated Cleaning System used for endoscope cleaning post an endoscopy procedure prior to high-level disinfection in the user's chosen Endoscope Washer Disinfector.

### Certified endoscopes

Contact your PFE Medical representative for information on certified endoscopes. If a new Endoscope is procured, please speak to your PFE Medical representative to ensure it is compatible and that the correct connectors are available before cleaning the Endoscope in the KHAMSIN.

- Double-channelled endoscopes will need to be cleaned in the KHAMSIN one at a time so that all the channels are connected e.g. Double Gastroscopes

- Endoscopes over 2.3 - 2.5m length with large channels will need to be cleaned in the KHAMSIN one at a time to ensure that the Endoscope is fully submerged e.g. Enteroscopes

- Any type of endoscope that has not been tested and trained on by PFE Training Staff will NOT have a barcode or RFID tag and can therefore not be processed in the KHAMSIN.

### Users and environment

The KHAMSIN is intended for use in hospitals and decontamination hubs.

It is designed to be used by healthcare professionals.

### Training

It is the responsibility of the department manager where the device is installed to ensure that:

      - All users are trained with regard to the instructions and warnings contained in this manual.

      - All users are aware of the risks and the preventive measures described in this manual.

# Recommended use

- The KHAMSIN is an automated cleaning system, replacing the manual cleaning process prior to the Endoscope being high-level disinfected as part of the endoscope reprocessing cycle.

- This device must be used in compliance with the operating procedure described in Section 4 "Use".

- The KHAMSIN chemical is to be used KHAMSIN with the KHAMSIN system to create a cleaning solution for endoscopes.

# Accessories

# Leak test connectors

# Endoscope connectors- amount may vary depending on endoscope make and model

# Button cleaning adaptors

# Technology

The KHAMSIN is an automated system using a combination of ultrasonics and nano-technology to clean both external and internal surfaces of flexible endoscopes.

# Certification

The KHAMSIN model is the KH01: it is compliant with Regulation (EU) 2017/745 of the European Parliament and of the Council of 5 April 2017 concerning Medical Devices and the Medical Devices Directive 93/42/EEC (UK Medical Device Regulations 2002 (SI 2002 No 618, as amended) (UK MDR 2002)).

The KHAMSIN is a class I device under these regulations.

This model carries the marking:

In accordance with the requirements of the above regulations, the KHAMSIN is compliant with the following standards:

| Standard | Application |
| --- | --- |
| EN 61326-1:2012 | Electromagnetic compatibility |
| IEC 61010-1:2010 IEC 61010-2-40:2015 | Electrical safety including specific requirements for sterilizers and disinfecting washers. Europe/US/CAN deviations |
| EN ISO 14971:2012 | Risk management |
| EN 62366-1:2015 | Application of usability engineering |
| IEC 62304:2015 | Medical-device software |
| EN 10993-1:2009 AC2010 | Biological evaluation of medical devices - Part 1: Evaluation and testing |

# Technical characteristics

| Dimensions | H78.5 cm x W75 cm x D77.5 cm |
| --- | --- |
| Weight | TBD |
| Power supply | 240v 3 pin plug |
| Power consumption | TBD |
| Disinfection efficacy | Cleaning only |
| Cleaning-cycle duration | 20 minutes |
| External operating temperature | +10°C to +40°C |
| Relative humidity | Less than 90% without condensation |
| Maximum operating altitude | 2000m |
| Acoustic power | 32 db |

## REGULATORY INFORMATION, SAFETY RULES, AND OPERATING PRECAUTIONS

# Safety rules

The following rules must be distributed by the department manager to all potential users of the KHAMSIN.

Please note:  Any serious incident that has occurred in relation to this device should be reported to the manufacturer and the Competent Authority.

### Handling

The User must never touch the ultrasonic pods at the bottom of the bowl enclosed within the body of the KHAMSIN when connected to power, to prevent risk of electrocution.

      - Do not use the KHAMSIN if there is any suspicion of damage.

      - Never force open the lid, especially not during an operating cycle.

      - Do not attempt to disassemble or repair the device. The KHAMSIN must be repaired exclusively by the manufacturer. Opening the equipment’s panels will result in voiding the warranty.

      - Do not bypass the equipment’s safety systems and interlocks.

      - Ensure the Endoscope is loaded into the cleaning chamber and connected correctly

### Environmental conditions

      - Do not store the KHAMSIN or its accessories at temperatures below -10°C or above 50°C.

### Use

      - The KHAMSIN is intended to be used exclusively for the use described previously, in the environment described previously, and by the personnel described previously. Do not use this equipment in any other way.

      - The KHAMSIN is not intended for the reprocessing of single-use devices.

      - The KHAMSIN is not a storage cabinet.

### Electricity

      - The equipment must be connected to a grounded power outlet. Ensure that the power cable used is effectively that provided with the equipment.

      - Connect the equipment only to a power supply having the appropriate voltage and frequency, as specified in Section 3.3 “Installation”. The use of an incorrect voltage may damage the equipment.

      - Any attempt to disassemble one of the components of the equipment may cause electrocution.

      - When shutdown is controlled by the switch on the control panel, certain parts of the cabling remain powered. The KHAMSIN must therefore be disconnected from the power supply before any technical intervention.

      - The user must take all precautions to avoid the ingress of liquid into the equipment other than for that intended for maintenance (see Section 5 “Maintenance”), which could create an electrical hazard.

### Electromagnetism

      - Warning: It is recommended to avoid using this equipment adjacent to other equipment or stacked on top of other equipment as this may result in incorrect operation. If this type of use is necessary, the operation of this equipment and any other equipment must be observed to verify normal operation.

      - Warning: Portable RF communication devices should not be used (including peripherals such as antenna cables and external antennas) at less than 30 cm (12 inches) from any part of the KHAMSIN, including the cables specified by the manufacturer. If not, the performance of such equipment may be adversely affected.

      - Warning: The use of accessories other than those specified or provided by the equipment manufacturer may result in an increase in electromagnetic emissions or a reduction of the equipment’s immunity and cause inappropriate operation.

### Upkeep and maintenance

Scheduled maintenance must be performed every year.

Non-compliance with this period and the recommended operations may result in dangerous situations as the equipment may be damaged, the user and third parties may be seriously injured and adjacent equipment and the environment may be at risk.

Only a technician approved by PFE Medical is authorised to perform maintenance operations on the KHAMSIN.

Scheduled maintenance includes:

- A general system check

- A cleaning of external and internal parts if necessary

- The replacement of wear parts, if necessary

- ACT test

- Verification of Calibrations

- Operational Checks

- Faults Verification

- Reproducibility of chemical additive

- Daily & weekly safety checks

- External faults

- Fault indications

- Chemical additive low level alarms

- Doors & interlocks

- Replacement of dosing pump tubing every 2 years or 5000 cycles

# Marking

| Model – KH01 Input – 230v AC AMPS - 8 max Frequency – 50 – 60 Serial number KH-XXX Manufacturer: Partners for Endoscopy | Rating plate: Manufacturer’s identification Description of the equipment Model Serial number Electrical characteristics Country of manufacture Markings UDI |
| --- | --- |
|  | Electric shock risk Only authorised personnel may open the access drawer to the maintenance zone. |
|  | Do not dispose of with municipal waste This symbol indicates that waste originating from electrical and electronic equipment (WEEE) must be handled with care and must not be disposed of with unsorted municipal waste. WEEE can pollute the environment and must be handled in a waste stream that is appropriate for WEEE. Contact the manufacturer or a company specialized in waste disposal to handle the disposal of this equipment. |
|  | Manufacturer This symbol adjacent to PFE Medical indicates that PFE Medical is the manufacturer of this equipment. |
|  | Information in the manual This symbol indicates that the user must read the manual before using the device. |
|  | Medical Device |
|  | EU Authorised Representative |
|  | UKCA |
|  | CE Mark |
|  | Unique Device Identifier |

## INSTALLATION, ADJUSTMENTS, AND COMMISSIONING

# Packaging contents

The packaging contains the following components:

| ITEMS | QUANTITY |
| --- | --- |
| KHAMSIN | 1 |
| MESH TRAY | 1 |
| SUCTION BIOPSY CONNECTOR (RED) | 2 |
| AIR/WATER CONNECTOR (BLUE) | 2 |
| AUXILIARY CONNECTOR (BLACK) | 2 |
| LEAK TEST CONNECTOR (YELLOW) | 2 |
| LEAK TEST BUNGS | 2 |
| PURPLE BLOCKER | 2 |
| BUTTON CLAMP | 2 |
| ENZYMATIC | (based on order qty) |
| CLEVER STUFF | (based on order qty) |
| BORING STUFF | (based on order qty) |
| CHEMICAL BAY INSERTS | (if required) |
| MANUAL | 1 |
| IFUS (LAMINATED) | 2 |
| DELIVERY NOTE | 2 |
| COSHH SHEETS 1 PER CHEMICAL | 1 |
| QR CODES VIA HOSPITAL PROVIDED SCOPE & USER LISTS | - |
| POT OF WIPES | 1 |
| PRINTER ROLL | 1 |
| TRAINING DOCUMENTS | - |

# Presentation of the interfaces

#### Connection to the AC power supply

The AC power connector is positioned on the KHAMSIN rear panel. Connection to AC power is via the IEC-standard cable provided with the equipment.

#### User interface

#### Touch Screen

#### The KHAMSIN has a touch screen which is for the user to start the KHAMSIN and to abort the cycle if needed when indicated. This also acts as a diagnostic display screen so that if the KHAMSIN does fail the screen will indicate the reason.

# Installation

Any damage during transport or storage may be a source of danger. Therefore, the user must check the integrity of the device before installation.

Verification

### Placement

The KHAMSIN needs a mixed water supply of potable water at between 32°- 37°C with ¾”BSP as well as a drain pipe with the diameter of 2”.  These need to be 400mm from ground level.

With a minimum 2bar dynamic water pressure The KHAMSIN also needs a 240V 3 pin socket outlet.

Please be aware - if the flow to the KHAMSIN is below our specification it will increase the cycle time.

Connection to electrical power

Warning

The Khamsin must be connected to an electrical installation that is compliant with the current regulations. This connection must be performed by a qualified person.

The KHAMSIN is connected to an AC power supply that must have a voltage of 240v. Check that the power source is in accordance with that indicated on the label. This product is of electrical class 2. Only power cables provided with the equipment may be used to connect it.

The power cable is used as a circuit-breaking system (cut-off of electrical power if necessary) and must therefore be accessible.

### Connection of accessories

Never connect peripherals other than those indicated below to the connectors.

- USE

# Switching on

As for any electrical equipment, malfunctions may arise over the lifetime of the product. Therefore, before switching on, ensure that the installation steps described above have been properly completed and that the maintenance has been performed in accordance with the indications set out in Section 6 “Scheduled maintenance”. In the event that these operations are not performed, the equipment’s interlocks could be undermined, the equipment may be damaged, the user and third parties may be seriously injured, and the environment may be at risk.

The device will be switched on using an ON/OFF switch located on the front of the device.

- To switch on the unit, turn the switch to on

- The devices’ LCD display panel will turn on

- The KHAMSIN is now powered on, meaning the device is ready for a cleaning cycle

# Switching off

The system should never be switched off directly using the AC power switch at the risk of irremediably damaging it and losing the saved traceability data.

- To switch off the unit, use the ON/OFF user switch on the front panel

- Then turn off at the socket

- The KHAMSIN is now powered off and has no power going to it

# Prerequisites to endoscope cleaning

Follow HTM guidance for the preparation of an endoscope prior to cleaning. At the time of this manual going to print the endoscope channels need to be flushed with detergent as soon as possible after the procedure prior to leaving the procedure room including a wipe of the external parts of the Endoscope with a bedside kit. (Cleanup)

# Cleaning an Endoscope Using the KHAMSIN

Install A MAXIMUM OF TWO endoscopes per cycle

Ensure that the endoscope does not come into contact with the equipment. Refer to section 1.1.2 on page 4

- The endoscope must be initially cleaned in the procedure room directly after the end of the procedure (Bedside clean-by drawing detergent through the working channel and cleaning the insertion tube ).

- To prevent protein drying on the endoscope, no more than 3 hours should pass between the procedure and the cleaning process.

Before using the KHAMSIN at the beginning of the list

- Check the display for any alerts for chemicals, and that print paper is present.

Cleaning an endoscope

- Ensure that the Endoscope has been cleaned at the bedside, directly after the procedure.

- Visually check the endoscope and load it into the chamber.**

The endoscope distal tip should be inserted into the mini central chamber in the centre of the main basin, especially those with a complicated distal tip (such as duodenoscopes or linear ultrasound endoscopes).

- Insert the suction valve/s into the U-clamps and place them into the chamber with the other Endoscope/s accessories. When reprocessing EBUS scopes, use a larger U clip for the suction button and a standard smaller U clip for the air /water button.

If you are reprocessing a duodenoscope with a removal sleeve, this must be placed in the mini chamber with the distal tip.

- Connect all relevant connectors: - blocker, air, water, suction, auxiliary jet, biopsy channel etc.

- Repeat steps 5-8 for the second endoscope, ensuring no part of the endoscope is sitting above the fluid line.

- Connect a leak tester. If only one scope is being reprocessed, insert the leak test blocker into the top leak test port.

- Visually check that all connectors are still correctly attached, remove dirty gloves and close the lid.

- Check the display panel that no alerts for chemistry are required; replenish appropriately.

- Press Login on the display.

- Using the User identification QR code/RFID tag, scan in the user.

- Select the number of endoscopes: 1 or 2.

- Using the QR code/RFID tag, scan in the first Endoscope.

- Press Log in again for a second scope.

- Using the QR code/RFID tag, scan in the second Endoscope.

- Press the start button.

- The KHAMSIN will carry out the automated leak test.

- Audibly confirm ultrasonics are active.

- At the end of the cycle, lift the lid and confirm all connections are still intact. If not, the process will need to be repeated.

- Remove the endoscope and continue with the decontamination process in the AER.

**In the case of extreme soiling, or the observation of “pips” or other elements e.g. worms, during the procedure, clean the Endoscope in a sink using the manual procedure as per your WPI’s before loading into the KHAMSIN.

IF THE ENDOSCOPE IS SUSPECTED TO HAVE BEEN IN CONTACT WITH vCJD, PROCESS THE ENDOSCOPE ON ITS OWN IN THE KHAMSIN AND THEN RUN AN EMPTY CYCLE AFTER THE ENDOSCOPE HAS BEEN REMOVED FROM THE KHAMSIN.

UPKEEP

## UPKEEP

# External cleaning

For cleaning the outside of the equipment, use a cloth or paper cloth, dampened with water or a surface cleaner/disinfectant. Machine Gleam is available for the external part of the KHAMSIN.

Do not spread liquid inside and over the KHAMSIN.

  - Cleaning of the chamber

For cleaning of the chamber use a surface disinfectant wipe. This should be done at the end of the day and more frequently if required.

## SCHEDULED MAINTENANCE

Scheduled maintenance must be performed every year.

Non-compliance with this period and the recommended operations may result in dangerous situations as the equipment’s interlocks could be undermined, the equipment may be damaged, the user and third parties may be seriously injure and adjacent equipment and the environment may be at risk.

Only a technician approved by PFE Medical is authorised to perform maintenance operations on the KHAMSIN.

Scheduled maintenance includes:

- A general system check

- A cleaning of external and internal parts if necessary

- The replacement of wear parts, if necessary

## TROUBLESHOOTING

The KHAMSIN is equipped with multiple sensors for proper operation.

If the KHAMSIN detects a malfunction during the cycle, even if only temporary, the cycle is stopped and the system goes into error mode. This function always guarantees the validity of the cleaning cycle.

PLEASE CALL YOUR PFE MEDICAL REPRESENTATIVE

## TRANSPORT AND STORAGE

Comply with the equipment’s storage and transport requirements

## DISPOSAL

When the KHAMSIN reaches the end of its life, PFE Medical will dispose of the equipment in accordance with the current applicable regulations.

## WARRANTY

The KHAMSIN has a 1-year warranty as of the unit’s date of delivery. This warranty covers defective parts and repair labor.

The KHAMSIN must be repaired exclusively by the manufacturer or a repair center authorised by the manufacturer. Contact your after-sales representative to organize this repair before returning the equipment.

## CONTACT

For any further information, contact your after-sales representative or the PFE Medical customer Service Department

PFE Medical customer service

Hartley House, Galveston Grove,

Fenton, Stoke on Trent, ST4 3PE, UK

sales@pfemedical.co.uk

Tel: +44 (0) 1782 594164