
INSTRUCTIONS FOR USE
EndoStream : Disposable water irrigation tubing for irrigation pumps (direct connection between the endoscope and sterile water bottle via an irrigation pump).
The EndoStream tubing is intended to be used for 24 hours and then disposed of. 
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PRODUCT CODES AND DESCRIPTIONS
	Product code
	Description
	Examples of compatible irrigation pumps


	PFE230 / SFP230 /

PFE230s / SFP230s
	Disposable water irrigation tubing for irrigation pumps
	EGP-100E®, Olympus OFP-1®, Erbe™

	SFP130/SFP130s
	Disposable water irrigation tubing for irrigation pumps
	Olympus® OFP-2, Pentax®, EndoStratus™, Scope Assist®

	PFE330 / SFP330 / 

PFE330s / SFP330s
	Universal Disposable water irrigation tubing 
	Olympus® OFP-2, Pentax®, EndoStratus™, Scope Assist®, EGP-100E®, Olympus OFP-1®, Erbe™


INTENDED USE 

The EndoStream tubing is designed to be used in conjunction with a peristaltic pump which attaches to a flexible endoscope to supply sterile water for irrigation during endoscopic procedures, on a patient indicated for an endoscopic procedure. The EndoStream tubing is compatible with Baxter, Versol and Fresenius Sterile water bottles.
EndoStream is a tubing with a cap which directly screws onto a sterile water bottle. The tubing extremity connects to the endoscope via the pump head on a peristaltic irrigation pump.
A back flow valve is integrated in the extremity of each EndoStream tubing ensuring the unidirectional flow from the sterile water bottle to the endoscope, mitigating the risks of cross contamination. 
INSTRUCTIONS FOR USE
Read all of the instructions for use for the disposable water irrigation tubing carefully before use. 
Please refer to the peristaltic pump operator’s manual for additional instructions for use
Consult the instructions for use of the endoscope as well as the concerned accessories before using this medical device.
This medical device must only be used by trained and qualified medical staff or authorized person familiar with endoscopy techniques and acquainted with the pathologies in question as well as their possible complications.

DIRECTIONS FOR USE
1) Open the sterile water bottle.
2) Remove EndoStream tubing from the sterile pouch.

3) Insert the tubing into the sterile water bottle and screw the plastic cap of the EndoStream to the sterile water bottle..

4) Peel off the 24h00 sticker from the packaging label. Write the date and time on the sticker before attaching it to the tubing.
5) Open the peristaltic pump head and place the tubing (PFE230/SFP130/SFP230/PFE330/SFP330 ensuring that it is placed between the markers).  
6) Close the peristaltic pump head
7) Connect the extremity of the EndoStream tubing to the auxiliary channel connector or to the lateral connexion of a biopsy valve. 

8) Prime the tubing before use by operating the peristaltic pump until sterile water flows through the digestive endoscope.
9) If the sterile water bottle should be changed during the procedure, consult trust guidelines on the correct technique in handling of the EndoStream during the replacement of the sterile water bottle, setup of the device and between patients.
10) When the procedure is completed, remove the EndoStream tubing, ensuring the blue check valve remains connected to the Endostream tubing.
11) Dispose of the EndoStream tubing after 24 hours.
WARNINGS AND PRECAUTIONS
· Always prime the auxiliary channel prior to inserting the digestive endoscope into the patient.
· The manufacturer did not design this device to be reprocessed or used for more than 24 hours once the packaging is opened, and therefore cannot verify that reprocessing can clean and/or sterilize or maintain the structural integrity of the device to ensure patient and/or user safety.  A traceability sticker is provided.
· Ensure the blue check valve is attached to the EndoStream at all times.
· The Endostream tubing is intended to be used for 24 hours after the packaging is opened. The product should be discarded within this time period regardless of numbers of use.

· Do not attach or reattach the EndoStream to an endoscope which has not been reprocessed.
· Only to be used with Sterile water bottles.

· Product must be discarded after it has been used in a procedure with vCJD.

· Do not use this product if the packaging is damaged.

Any serious incident that has occurred in relation to this device should be reported to the manufacturer and Competent Authority.
CONTRAINDICATIONS
The contraindications are those which are specific to all endoscopic procedures.
The disposable irrigation tubings should not be re-sterilized.
STORAGE / PACKAGING
Storage : In order to prevent any damage, products should be stored in their original packaging, away from sunlight in a cool dry place.
Packaging : The disposable irrigation tubings, EndoStream, are supplied in an individual sterile package in a box of 10 units.
PRODUCT DISPOSAL
After use, this product may be a potential biohazard. Handle and dispose of in accordance with accepted medical practice.
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Advena Limited, Tower Business Centre, 2nd Floor, Tower Street, Swatar, BKR 4013, Malta 
Instructions for use V22 July 2026, Olympus® is a registered trademark of Olympus Optical Co., Ltd. Pentax® is a registered trademark of Pentax Medical
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